Standard Operating Procedure

| Procedure Title: | # Auto Clave EP |
| --- | --- |
| Date: | 10/07/2024 |
| Version: | 1.0 |
| Department: | Primary Care EP |
| Approved By: | Cece Carvajal |

### 1. Purpose

This SOP provides standardized instructions for safely operating the autoclave to ensure proper sterilization of instruments and materials.

### 2. Scope

This procedure applies to all staff responsible for sterilization using the autoclave.

### 3. Responsibilities

- Staff are responsible for correctly following all steps.

- Supervisors are responsible for ensuring training and compliance.

### 4. Materials Needed

- Autoclave

- Distilled water

- Items to be sterilized (in approved sterilization packaging/trays)5. Procedure

#### A. Preparing the Autoclave

- Fill Reservoir:

  - Pour distilled water into the reservoir through the opening on the top right of the autoclave.

  - Fill until water reaches:

    - The base of the safety valves, and

    - The indicator grooves in the front of the chamber.

- Fill Chamber:

  - Turn the control knob to FILL position.

  - Allow water to enter the chamber until it reaches the indicator groove.

- Set to Sterilize:

  - Turn the control knob to STERILIZE position.

#### B. Loading and Sterilization

- Load Items: Place instruments/materials into the autoclave chamber.

- Close Door: Ensure the autoclave door is tightly closed.

- Set Parameters:

  - Temperature: 273°F

  - Timer: Set to required sterilization time 30 minutes

#### C. End of Sterilization Cycle

- When the timer reaches zero:

  - Turn the control knob to EXH + DRY.

- Liquids Only:

  - Leave the valve in the same sterilization position at cycle completion.

  - Allow pressure to decrease slowly.

  - Once pressure reaches zero, carefully open the autoclave door.

  - Note: This may take 30 seconds to 1 minute.

#### D. Drying Process

- After sterilization:

  - Open the autoclave door approximately 1 inch.

  - Add 20–30 minutes to the timer for drying.

- Once drying cycle is complete:

- Turn the control knob from EXH + DRY back to 0 (OFF).

### 6. Safety Notes

- Never open the autoclave door until pressure gauge reads zero.

- Use heat-resistant gloves when handling items.

- Always use distilled water (tap water can damage the unit).